## Sample Submission
[7 empty] | last updated | 2022-11-01
UWPR sample submission form for LC-MS/MS analyses:
Only samples ready for mass spectrometry analysis are accepted. At this time we do not offer any sample preparation.
Radioactive samples are not allowed!
Before submitting any samples:
à Please contact us to discuss your project.
à Create a project and submit billing information online and make sure you select “Mass Spec. analysis by UWPR personnel?”:
https://proteomicsresource.washington.edu/
Then complete this form and upload it to your project online and email priska@uw.edu to coordinate a time to drop off your samples.
Once we receive this form and your samples we schedule the instrument time needed for the analysis. Note: cancellation time is 48hrs, 10% (of total cost) sign up fee is non-refundable.
* | Customer Information (all fields required)
Principal Investigator (PI):
| PI is aware of and approved this order? |  |  |  | Submission Date
| Submitted by |  |  |  | Project #
| Phone |  |  |  | Budget # |  |  | Make sure the budget number(s) are entered online under your project and that they are current. You can use up to two budget numbers to split the cost, please specify the percentage for each number.
| Email address: |  |  |  | Number of samples | 1
[5 empty] | Number of blanks | 0
Sample information (all fields required)
* | Brief description:
* | Species to use for data analysis:
* | Sample label (as it appears on vials) in the order you want them analyzed: |  | e.g. sample1, blank, sample2, QC, sample3 ….
fill in as much information belwo as you know or any special instructions for the analysis:
| Sample’s Return Requested? |  |  | samples unclaimed within two weeks of analysis will be discarded, sample storage at -20oC
| Sample Format (please select): |  |  | if liquid list buffer composition below
Sample Composition (salts, buffers, pH etc.):
| Concentration (µg or µg/µl) |  |  | µg/µl |  | µg (if dry) | complex peptide mixture: 0.1 - 0.5µg in 1 - 8µl on column
| Injection volume per LC-MS run |  |  | µl | range 1-8 µl |  | single peptide 0.1-2 pmol in 1 - 8 µl on column
Method of Purification
Reduction & Alkylation:
Modifications:
Modification detail (e.g. 10-plex etc.):
Amino Acids Modified:
Digestion Enzyme (other than Trypsin):
Analysis details:
Column information:
| Analytical column ID 75 µm, beads Dr. Maisch ReprosilPur C18AQ 5µm 120Å, length 35 cm | [5 empty] | cm (specify custom length if desired)
| Trap column ID 100 µm, beads Dr. Maisch ReprosilPur C18AQ 5µm 120Å, length 3 cm | [5 empty] | cm (specify custom length if desired)
Gradient information:
Buffer A: 0.1% Formic Acid in Water (Optima™ LC/MS, Solvent Blends, Fisher Chemical)
Buffer B: 0.1% Formic Acid in 80% Acetonitrile (Optima™ LC/MS, Solvent Blends, Fisher Chemical)
Trapping default: 18µl, 100% A at 2.5 µl/min
Gradient default (Medium): 6 - 35% B in 90 mins; 100% B for 10min; 2% B for 30 mins, 10 min trapping; total analysis time = ~140 mins per LC-MS run
To specify different gradient list here and check the checkbox below:
Enter the number of samples, replicate analyses per sample, blanks and QC's to calculate the time needed for the analysis:
| Number of samples: | 1 |  |  |  | Number of replicate LC-MS analyses per sample: | 1
| Number of blanks: | 0 |  |  |  | Number of extra QC (standards): | 0
| Analytical Gradients: | Analysis time [min] | # analyses |  | Short | Medium | Long | Extra long | Custom
| Short [60 mins] | 110 | 1 |  | 2 |  |  |  |  | Short: 6 - 35% B in 60 mins; 80% B for 10min; 2% B for 30 mins, 10 min trapping à total analysis time = 110 mins per LC-MS run
| Medium [90 mins] (default) | 140 | 1 |  |  | 3 |  |  |  | Medium: 6 - 35% B in 90 mins; 80% B for 10min; 2% B for 30 mins, 10 min trapping à total analysis time = 140 mins per LC-MS run
| Long [120 mins] | 170 | 1 |  |  |  | 3 |  |  | Long: 6 - 35% B in 120 mins; 80% B for 10min; 2% B for 30 mins, 10 min trapping à total analysis time = 170 mins per LC-MS run
| Extra long [180 mins] | 230 | 1 |  |  |  |  | 4 |  | Extra long: 6 - 35% B in 180 mins; 80% B for 10min; 2% B for 30 mins, 10 min trapping à total analysis time = 230 mins per LC-MS run
| Custom specify your own: |  | 1 | [5 empty] | 0 | Custom specify your own:
| QC (quality control) | 75 | 0 |  | 0 | 0 | 0 | 0 | 0 | QC (AngioNeuro std) runs, analysis time is 75 min per QC
| blanks | 40 | 0 |  | 0 | 0 | 0 | 0 | 0 | min run time per blank is 40 min
|  |  | Est. total hrs: |  | 4 | 5 | 5 | 6 | 2 | incl. 2 hrs minimum setup (colunn preparation, hookup, equilibration and 1 QC)
|  |  | Cost Estimate: |  | 149 | 180 | 180 | 211 | 87
Non-refundable sign up fee of 10% is applied when instrument time is scheduled.
The remaining 90% of the total cost will be applied only if instrument time is used, i.e. not canceled 48 hrs prior to start time.
Note this is just an estimate based on the information you provided in this document.
The actual cost will be determined after your samples are analyzed and reported in our online system.
UW Internal Rates With Additional Labor
| Select Instrument | [11 empty] | TSQA | TSQV | Exploris 1 | Exploris 2 | QE + | Fusion | Lumos
[12 empty] | Hourly rate | 27 | 18 | 31 | 31 | 31 | 31 | 31
| MS acquisition | if left blank we use UWPR default method settings | [9 empty] | Setup fee | 25 | 25 | 25 | 25 | 25 | 25 | 25
| Full scan | m/z range: |  |  | Resolution: | [5 empty] | Gradient:
| MS/MS scan | Top N/cyle time: | [8 empty] | Short |  | 133 | 97 | 149 | 149 | 149 | 149 | 149
|  | m/z range: |  |  | Resolution: | [5 empty] | Medium |  | 160 | 115 | 180 | 180 | 180 | 180 | 180
| Specify and additional settings or scans: | [9 empty] | Long |  | 160 | 115 | 180 | 180 | 180 | 180 | 180
[11 empty] | Extra long |  | 187 | 133 | 211 | 211 | 211 | 211 | 211
[11 empty] | Custom |  | 79 | 61 | 87 | 87 | 87 | 87 | 87
| Database Search Information | [9 empty] | Non-refundable sign up fee of 10% is applied when instrument time is scheduled via the web portal
| Database search performed by UWPR? | [9 empty] | The remaining 90% of the total cost will be applied only if instrument time is used, i.e. not canceled
| Database search includes Comet search, and Peptide/Protein profit analysis, please discuss details with Jimmy engj@uw.edu | [9 empty] | Note this is just an estimate based on the information you provided in this document.
| Species to use for data analysis: |  | 0 | [7 empty] | The actual cost will be determined after your samples are analyzed and reported in our online system.
If desired, select modifications to be searched below
describe labels (10 plex, Lys label… ):
| Specify amino acid(s) and mass: |  |  | aa |  | mass (monoisotopic)
yes
no
Dry/Solid
Liquid
Fusion
Q Exactive
Exploris
centroid
profile
centroid
profile
CID
HCD
ETD
yes
no
oxidation (M)
carboxymethyl (C) if using iodoacetic acid
carbamidomethyl (C) if using iodoacetamide
iTRAQ
TMT
SILAC
other (please specify below)
Lumos
phosphorylation
other
iTRAQ
TMT
SILAC
yes
no
yes
no
Note on A13: required
Note on B15: NOTE: your PI will receive an email when UWPR personnel schedules instrument time
Note on G17: enter budget number for your project online and make sure the budget number online is current
Note on A22: required
Note on A23: required
Note on A24: required

## methods
Methods | (to be completed by UWPR)
Mass Spec method |  | HPLC method
Method drafts (Waters nanoAcquity)
For detailed method information specific to your experiment open the raw file in Qual Browser. Highlight the green pushpin and do a right mouse click, select View/Report/Instrument method
Use the left and right arrow to toggle between mass spec and HPLC method
To copy the text:
Ctrl A and ctrl C, then go to word or xls and ctrl V
Use the left and right arrow to toggle between mass spec and HPLC method
Below is a draft for your methods, make sure you check and change all the details based on your specific experiment.
Mass Spectrometry Analysis Fusion/Lumos
All samples were analyzed on an Orbitrap Fusion Lumos mass spectrometer (ThermoFisherScientific) equipped with a nano-Acquity UPLC system (Waters) and in house developed nano spray ionization source.
Samples (xx μl at xx μg/μl) were loaded from the autosampler onto a 100 μm ID Integrafrit trap (NewObjective) packed with Reprosil-Pur C18-AQ 120 Å 5 µm material (Dr. Maisch) to a bed length of 2.5 cm at a
flow rate of 2 µL/min. After loading and desalting for 10 min with 0.1% formic acid plus 2% acetonitrile (LCMS grade from Fisher), the trap was brought in-line with a pulled fused-silica capillary tip (75-μm i.d.)
packed with 35 cm of Reprosil-Pur C18-AQ 120 Å 5 µm (Dr. Maisch). Peptides were separated using a linear gradient, from 5-30% solvent B (LCMS grade 0.1 % formic acid it acetonitrile (Fisher)) in 90 min at a
flow rate of 300 nL/min. The column temperature was maintained at a constant 50oC during all experiments.
Peptides were detected using a data-dependent (DDA) method. Survey scans of peptide precursors were performed in the orbitrap mass analyzer from 375 to 1575 m/z at 120K resolution (at 200 m/z) with a
7e5 ion count target and a maximum injection time of 50 ms. The instrument was set to run in top speed mode with 3 sec cycle for the survey and the MS/MS scans.
IT MS2 (if msms scans were acquired in the ion trap)
After a survey scan, tandem MS was performed on the most abundant precursors exhibiting a charge state from 2 to 5 of greater than 1e4 intensity by isolating them in the quadrupole with an isolation width of 1.6 m/z.
Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 30 % and resulting fragments were detected using the rapid scan rate in the ion trap. The AGC target
for MS/MS was set to 2e3 and the maximum injection time limited to 100 ms. The dynamic exclusion was set to 30 sec with a 20 ppm mass tolerance around the precursor and its isotopes. Monoisotopic precursor
selection was enabled.
OT MS2 (if msms scans were acquired in the orbitrap)
After a survey scan, tandem MS was performed on the most abundant precursors exhibiting a charge state from 2 to 5 of greater than 2e4 intensity by isolating them in the quadrupole with an isolation width of 1.6 m/z.
Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 30 %. Resulting fragments were detected in the orbitrap mass analyzer at 30K resolution (at 200 m/z) with a
1e4 ion count target and a maximum injection time of 100 ms. The dynamic exclusion was set to 30 sec with a 20 ppm mass tolerance around the precursor and its isotopes. Monoisotopic precursor selection was enabled.
Mass Spectrometry Analysis Q Exactive Plus
All samples were analyzed on an Q Exactive Plus mass spectrometer (ThermoFisherScientific) equipped with a nano-Acquity UPLC system (Waters) and in house developed nano spray ionization source.
Samples (xx μl at xx μg/μl) were loaded from the autosampler onto a 100 μm ID Integrafrit trap (NewObjective) packed with Reprosil-Pur C18-AQ 120 Å 5 µm material (Dr. Maisch) to a bed length of 2.5 cm at a
flow rate of 2 µL/min. After loading and desalting for 10 min with 0.1% formic acid plus 2% acetonitrile (LCMS grade from Fisher), the trap was brought in-line with a pulled fused-silica capillary tip (75-μm i.d.)
packed with 35 cm of Reprosil-Pur C18-AQ 120 Å 5 µm (Dr. Maisch). Peptides were separated using a linear gradient, from 5-30% solvent B (LCMS grade 0.1 % formic acid it acetonitrile (Fisher)) in 90 min at a
flow rate of 300 nL/min. The column temperature was maintained at a constant 50oC during all experiments.
Peptides were detected using a data-dependent (DDA) method. Survey scans of peptide precursors were performed in the orbitrap mass analyzer from 400 to 2000 m/z at 70K resolution (at 400 m/z) with
a 1e6 ion count target and a maximum injection time of 100 ms. After the survey scan, tandem MS was performed on the top 20 most abundant precursors exhibiting a charge state from 2 to 5 of greater
than 1e4 intensity by isolating them in the quadrupole with an isolation width of 1.6 m/z. Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 25 %.
Resulting fragments were detected in the orbitrap mass analyzer at 17.5K resolution (at 400 m/z) with a 5e4 ion count target and a maximum injection time of 50 ms. The dynamic exclusion was set to 30 sec
and isotopes were excluded.
Method drafts (Thermo EASYnLC)
For detailed method information specific to your experiment open the raw file in Qual Browser. Highlight the green pushpin and do a right mouse click, select View/Report/Instrument method
Use the left and right arrow to toggle between mass spec and HPLC method
In Freestyle open the RAW file click on instrument method and select the instrument or HPLC to display the method
To copy the text:
Ctrl A and ctrl C, then go to word or xls and ctrl V
Use the left and right arrow to toggle between mass spec and HPLC method
Below is a draft for your methods, make sure you check and change all the details based on your specific experiment.
Mass Spectrometry Analysis Fusion/Lumos
All samples were analyzed on an Orbitrap Fusion Lumos mass spectrometer (ThermoFisherScientific) equipped with an EASYnLC 1200 UPLC system (ThermoFisherScientific) and in house developed nano spray ionization source.
Samples (xx μl at xx μg/μl) were loaded from the autosampler onto a 100 μm ID Integrafrit trap (NewObjective) packed with Reprosil-Pur C18-AQ 120 Å 5 µm material (Dr. Maisch) to a bed length of 3 cm with a volume of 18 µl at a
flow rate of 2.5 µL/min. After loading and desalting with 0.1% formic acid in water (LCMS grade from Fisher), the trap was brought in-line with a pulled fused-silica capillary tip (75-μm i.d.)
packed with 35 cm of Reprosil-Pur C18-AQ 120 Å 5 µm (Dr. Maisch). Peptides were separated using a linear gradient, from 6-35 % solvent B (LCMS grade 0.1 % formic acid, 80 % acetonitrile in water (Fisher)) in 90 min at a
flow rate of 300 nL/min. The column temperature was maintained at a constant 50oC during all experiments.
Peptides were detected using a data-dependent (DDA) method. Survey scans of peptide precursors were performed in the orbitrap mass analyzer from 375 to 1575 m/z at 120K resolution (at 200 m/z) with a
7e5 ion count target and a maximum injection time of 50 ms. The instrument was set to run in top speed mode with 3 sec cycle for the survey and the MS/MS scans.
DDA using IT MS2 (if msms scans were acquired in the ion trap)
After a survey scan, tandem MS was performed on the most abundant precursors exhibiting a charge state from 2 to 5 of greater than 1e4 intensity by isolating them in the quadrupole with an isolation width of 1.6 m/z.
Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 30 % and resulting fragments were detected using the rapid scan rate in the ion trap. The AGC target
for MS/MS was set to 2e3 and the maximum injection time limited to 100 ms. The dynamic exclusion was set to 30 sec with a 20 ppm mass tolerance around the precursor and its isotopes. Monoisotopic precursor
selection was enabled.
DDA using OT MS2 (if msms scans were acquired in the orbitrap)
After a survey scan, tandem MS was performed on the most abundant precursors exhibiting a charge state from 2 to 5 of greater than 2e4 intensity by isolating them in the quadrupole with an isolation width of 1.6 m/z.
Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 30 %. Resulting fragments were detected in the orbitrap mass analyzer at 30K resolution (at 200 m/z) with a
1e4 ion count target and a maximum injection time of 100 ms. The dynamic exclusion was set to 30 sec with a 20 ppm mass tolerance around the precursor and its isotopes. Monoisotopic precursor selection was enabled.
Parallel Reaction Monitoring (PRM)
Peptides were detected using a targeted Parallel Reaction Monitoring (PRM) method. After the survey scan, targeted MSMS was performed based on the inclusion list of N precursors (m/z, charge state) generated by Skyline.
Precursors were isolated in the quadrupole with an isolation width of 1.6 m/z.
Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 30 % and resulting fragments were detected in the orbitrap mass analyzer at 30K resolution (at 200 m/z) with a
100% normalized AGC target and a maximum injection time of 100 ms. The loop count was set to 'All' or N=xx, to generate xx fragment ion spectra per MS1 scan.
Mass Spectrometry Analysis Exploris480
All samples were analyzed on an Exploris480 mass spectrometer (ThermoFisherScientific) equipped with an EASYnLC 1200 UPLC system (ThermoFisherScientific) and in house developed nano spray ionization source.
Samples (xx μl at xx μg/μl) were loaded from the autosampler onto a 100 μm ID Integrafrit trap (NewObjective) packed with Reprosil-Pur C18-AQ 120 Å 5 µm material (Dr. Maisch) to a bed length of 3 cm with a volume of 18 µl at a
flow rate of 2.5 µL/min. After loading and desalting with 0.1% formic acid in water (LCMS grade from Fisher), the trap was brought in-line with a pulled fused-silica capillary tip (75-μm i.d.)
packed with 35 cm of Reprosil-Pur C18-AQ 120 Å 5 µm (Dr. Maisch). Peptides were separated using a linear gradient, from 6-35 % solvent B (LCMS grade 0.1 % formic acid, 80 % acetonitrile in water (Fisher)) in 90 min at a
flow rate of 300 nL/min. The column temperature was maintained at a constant 50oC during all experiments.
DDA
Peptides were detected using a data-dependent (DDA) method. Survey scans of peptide precursors were performed in the orbitrap mass analyzer from 380 to 1575 m/z at 120K resolution (at 400 m/z) with
300% normalized AGC target and a maximum injection time of 25 ms. After the survey scan, tandem MS was performed for 1 second on the most abundant precursors exhibiting a charge state from 2 to 5 of greater
than 1e4 intensity by isolating them in the quadrupole with an isolation width of 1.6 m/z. Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 30 %.
Resulting fragments were detected in the orbitrap mass analyzer at 15K resolution (at 400 m/z) with a 100% normalized AGC target and a maximum injection time of 40 ms. The dynamic exclusion was set to 30 sec
and isotopes were excluded.
Parallel Reaction Monitoring (PRM)
Peptides were detected using a targeted Parallel Reaction Monitoring (PRM) method. After the survey scan, targeted MSMS was performed based on the inclusion list of N precursors (m/z, charge state) generated by Skyline.
Precursors were isolated in the quadrupole with an isolation width of 1.6 m/z.
Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 30 % and resulting fragments were detected in the orbitrap mass analyzer at 15K resolution (at 400 m/z) with
a 100% normalized AGC target and a maximum injection time of 100 ms. The loop count was set to 'All' or N=xx, to generate xx fragment ion spectra per MS1 scan.
Mass Spectrometry Analysis Q Exactive Plus
All samples were analyzed on an Q Exactive Plus mass spectrometer (ThermoFisherScientific) equipped with an EASYnLC 1200 UPLC system (ThermoFisherScientific) and in house developed nano spray ionization source.
Samples (xx μl at xx μg/μl) were loaded from the autosampler onto a 100 μm ID Integrafrit trap (NewObjective) packed with Reprosil-Pur C18-AQ 120 Å 5 µm material (Dr. Maisch) to a bed length of 3 cm with a volume of 18 µl at a
flow rate of 2.5 µL/min. After loading and desalting with 0.1% formic acid in water (LCMS grade from Fisher), the trap was brought in-line with a pulled fused-silica capillary tip (75-μm i.d.)
packed with 35 cm of Reprosil-Pur C18-AQ 120 Å 5 µm (Dr. Maisch). Peptides were separated using a linear gradient, from 6-35 % solvent B (LCMS grade 0.1 % formic acid, 80 % acetonitrile in water (Fisher)) in 90 min at a
flow rate of 300 nL/min. The column temperature was maintained at a constant 50oC during all experiments.
DDA
Peptides were detected using a data-dependent (DDA) method. Survey scans of peptide precursors were performed in the orbitrap mass analyzer from 400 to 2000 m/z at 70K resolution (at 400 m/z) with
a 1e6 ion count target and a maximum injection time of 100 ms. After the survey scan, tandem MS was performed on the top 20 most abundant precursors exhibiting a charge state from 2 to 5 of greater
than 1e4 intensity by isolating them in the quadrupole with an isolation width of 1.6 m/z. Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 25 %.
Resulting fragments were detected in the orbitrap mass analyzer at 17.5K resolution (at 400 m/z) with a 5e4 ion count target and a maximum injection time of 50 ms. The dynamic exclusion was set to 30 sec
and isotopes were excluded.
Parallel Reaction Monitoring (PRM)
Peptides were detected using a targeted Parallel Reaction Monitoring (PRM) method. After the survey scan, targeted MSMS was performed based on the inclusion list of N precursors (m/z, charge state) generated by Skyline.
Precursors were isolated in the quadrupole with an isolation width of 1.6 m/z.
Higher-energy collisional dissociation (HCD) fragmentation was applied with a normalized collision energy of 30 % and resulting fragments were detected in the orbitrap mass analyzer at 30K resolution (at 200 m/z) with a
1e4 ion count target and a maximum injection time of 100 ms. The loop count was set to 'All' or N=xx, to generate xx fragment ion spectra per MS1 scan.
For database search details check with Jimmy, here is a draft:
Thermo .raw files were converted to the mzXML format using the ReAdW (version 2016.1.0) converter. The mzXML files were searched against a Species RAT.fasta protein sequence database downloaded from UniProt
appended with common contaminant sequences (24,000 total sequence entries). A database search was performed using Comet ( 2021.02 rev. 0 (3c62af2) ) with the following search parameters: 20 ppm precursor tolerance,
concatenated target-decoy search, tryptic digest allowing 2 missed cleavages, variable modifications for oxidized methionine and phosphorylation on STY, and static modification for carboxyamidomethylation on cysteine.
The Comet search results were then processed with PeptideProphet and ProteinProphet tools from the Trans-Proteomic Pipeline software suite ( TPP v5.0.0 Typhoon ).
Acknowledgements:
For publications, the Proteomics Resource should always be acknowledged as follows in publications:
This work is supported in part by the University of Washington's Proteomics Resource (UWPR95794).

## sequence
to be completed by UWPR | [5 empty] | start | end
File Name | Comment | Path | Instrument Method | Inj Vol | Position

## screenshots
TIC of samples | TIC of QC before samples
TIC of QC after samples

## Rates
Rates effective 11/1/2022 subject to change without notice | [5 empty] | Revised 10/28/2022 |  |  | Instrumentation:
[9 empty] | Exploris 1 | Orbitrap Exploris
[9 empty] | Exploris 2 | Orbitrap Exploris
[9 empty] | Lumos | Orbitrap Fusion Lumos
[9 empty] | Fusion | Orbitrap Fusion
[9 empty] | QE+ | Q Exactive plus
[9 empty] | TSQA | TSQ Altis
[9 empty] | TSQV | TSQ Vantage
for a detailed description check out our resources webpage
UW Internal Rates Without Labor
|  |  | TSQA | TSQV | Exploris 1 | Exploris 2 | QE + | Fusion | Lumos
| Hourly Rate |  | 13.5 | 9 | 15.5 | 15.5 | 15.5 | 15.5 | 15.5
| Enter the number of hours here: | 1
| Total instrument time: |  | 13.5 | 9 | 15.5 | 15.5 | 15.5 | 15.5 | 15.5
| Setup fee per consecutive time block |  | 25 | 25 | 25 | 25 | 25 | 25 | 25
| Total Cost: |  | 38.5 | 34 | 40.5 | 40.5 | 40.5 | 40.5 | 40.5
UW Internal Rates With Additional Labor
|  |  | TSQA | TSQV | Exploris 1 | Exploris 2 | QE + | Fusion | Lumos
| Hourly Rate |  | 27 | 18 | 31 | 31 | 31 | 31 | 31
| Enter the number of hours here: | 1
| Total instrument time: |  | 27 | 18 | 31 | 31 | 31 | 31 | 31
| Setup fee per consecutive time block |  | 25 | 25 | 25 | 25 | 25 | 25 | 25
| Total Cost: |  | 52 | 43 | 56 | 56 | 56 | 56 | 56
Start
Mon-Fri
End
24/7
Start
Mon-Fri
6 pm
2 pm
10 am
- Instrument time can start either 10am or 2pm Mon-Fri.
- End times can be 10am, 2pm and 6pm 7 days a week.
- Consecutive instrument time is charged based on hourly rates
plus one setup charge
- Cancellation time is 48hrs prior to start time

## Concentrations
Recommended concentrations
In-solution samples | At least several hundred femtomole of protein in 10-20 µL of sample
Injection volume 1-5 µl is ideal, up to 8µl max
For complex mixtures up to 0.5µg (QE/Fusion/Lumos)
MRM on TSQVantage: some peptides can be quantified in the low amol range (~10 amol) others require more like low fmol
General Advice:
Samples can be submitted in solution (min 10-20 µL volume). Buffer should be 0.1% formic acid in 0 - 5% acetonilrile in water. Avoid any detergents and unnecessary sample buffers. For a comprehensive list of buffers and salts compatible with mass spectrometry, please see the Salt Tolerance Table. A BCA or Bradford assay (or similar) should be done to estimate the amount of protein in solution before submission and the amount indicated on the form. Ideally, we would like several hundred femtomoles to 1 pmol of protein for identification. If you need to clean up your sample, you can find some tips on our website:
http://www.proteomicsresource.washington.edu/protocols03/
In-gel-digest samples | Start with min of one picomole of protein (in gel)
General Advice about in-gel digestion:
It is very kritical to avoid contaminations, particularly keratin from skin and hair. Work clean, wipe all surfaces that come in contact with the gel e.g. with ethanol and wear gloves at all times.
There are many protocols out there, including on our website.
http://www.proteomicsresource.washington.edu/protocols03/ingeldigestion.php
https://tools.thermofisher.com/content/sfs/brochures/TR0050-Stained-gels-for-MS.pdf
We recommend gels be stained with either Commassie Blue or Sypro Ruby. These produce the best results. Stains with colloidal coomassie G-250 generally produce “better stains”, most common commassie blue stains are capatible with mass spectrometry. Silver stained gels generally produce less robust results because they modify the proteins and crosslink the peptides to the gels. If you need to use silver stains, they must be mass spec compatible. Most commercial suppliers will indicate this on their product insert.
| e.g. | Pierce Silver Stain Kit for Mass Spectrometry
Avoid contaminations: | Contaminations compete with your peptides of interest!
Contaminants bind to your column, preventing peptides to bind
Contaminants will be analyzed by the mass spec, the more MS2 scans are wasted on contaminants, the fewer MS2 scans are available to identify your peptides
Contaminants can clog columns costing you instrument time.
Contaminants can reduce instrument sensitivity, we will charge you for instrument cleaning minimum of 2 days or 48hrs.
Salt tolerances: | no salts! |  | salt tolerances
Detergent tolerances: | no detergents! |  | detergent tolerances
Lipids: | no lipids!
Other ources for contamination |  |  | glassware previously exposed to detergents
Avoid Contaminations (pdf) |  |  | plastics
acids contaminated with plastic from pipett tips
contaminations |  |  | Detergents (e.g. Triton, SDS)
PEG
Polymers
Sephadex
DTT
BME
HEPES (or other non-volatile buffer)
EDTA
Phosphate salts
Flag tag elution
Elution reagents
Protease inhibitors (e.g. AEBSF, Roche Mini Complete, aprotinin, leupeptin)
Phosphatase inhibitors
Any enzymes
Reduction and alkylation
Denaturants (e.g. urea)

## Salt Tolerances
This table for Salt Tolaraces is originally from the Scripps website, but we are adding new concentrations as they become available.:
http://masspec.scripps.edu/services/proteomics/images/saltbuffer.pdf
Surfactant, Buffer and Salt | Mw (g/mol.) | MALDI (mM) | MALDI (wt.%) | ESI (mM) | ESI (wt.%) | Reference*
BICINE | 163 | 50 | 0.8 | n. a. | n. a. | B
Brij |  |  |  |  | 0.03 | G
CHAPS | 615 | 0 | 0 | 1.6 | 0.1 | G,C, E
CTAB | 284 | n. a. | n. a. | <3.5 | <0.1 | F
Dithiothreitol | 154 | 500 | 7.7 | n. a. | n. a. | D
Glycerol | 92 | 130 | 1.2 | n. a. | n. a. | C, D
Guanidine, HCl | 96 | 250 | 2.4 | n. a. | n. a. | C, D
HEPES | 238 | 100 | 2.4 | n. a. | n. a. | A, B
IGEPAL CA-630 |  |  | 0 |  | 0.01
LDAO | 229 | 4.4 | 1 | <4.4 | <0.1 | C, F
n- Hexyl glucoside | 264 | n. a. | n. a. | 3.8 | 0.1 | E
NaCl | 58 | 50 | 0.29 | n. a. | n. a. | C, D
NaHPO4 | 120 | 10 | 0.12 | 10 | 0.12 | B, C, D, F
n-Dodecyl glucoside | 348 | n. a. | n. a. | 2.9 | 0.1 | E
n-Dodecyl maltoside | 511 | n. a. | n. a. | 2 | 0.1 | E
n-Dodecyl sucrose | 524 | n. a. | n. a. | 1.9 | 0.1 | E
NH4HCO3 | 79 | 50 | 0.4 | n. a. | n. a. | D
N-Octyl- -glucopyranoside | 292 | 3.4 | 0.1 | 3.4 | 0.1 | C, E,G
n-Octyl sucrose | 468 | n. a. | n. a. | 2.1 | 0.1 | E
NP-40 | 603 | 1.7 | 0.1 | n.a. | n. a.
Octyl thioglucoside | 308 | n. a. | n. a. | 3.2 | 0.1 | E
PEG1000 | 1000 | n. a. | n. a. | 0.5 | 0.05 | F
PEG2000 | 2000 | 0.5 | 0.1 | n. a. | n. a. | C,
SDS | 288 | 0.35 | 0.01 | 0.335 | 0.01 | C, D, E, F
Sodium Acetate | 82 | 50 | 0.41 | n. a. | n. a. | B, C
Sodium Azide | 65 | 15 | 0.1 | 3.1 | 0.02 | C, F
Sodium Cholate | 431 | n. a. | n. a. | 2.3 | 0.1 | E
Sodium Taurocholate | 538 | n. a. | n. a. | <1.9 | <0.1 | F
TFA | 114 | n. a. | n. a. | 4.4 | 0.05 | Pri. Comm.
Thesit | 583 | n. a. | n. a. | <1.7 | <0.1 | E
TRIS | 121 | 100 | 1 | n. a. | n. a. | A, B
Triton X-100, | 628 | 1.6 | 0.1 | <1.6 | 0.05 | C, E, G
Tween20 | 1228 | n. a. | 0.6 | n.a. | 0.05 | E, G
Urea | 60 | 500 | 3 | n. a. | n. a. | C, D
Zwittergent, 3-16 | 392 | 2.6 | 0.1 | n. a. | n. a. | C
*Note: When multiple recommendations were available, the lower (or lowest)
concentration was chosen.
HEPES: N-[2-hydroxyethyl]piperazine-N’-[2-ethanesulfonic acid]
TRIS: Tris[hydroxymethyl]aminomethane acetate
BICINE: N,N-bis[2-hydroxyethyl]glycine
CHAPS: 3-[(3-cholamidopropyl)-dimethylammonio]-1-propane sulfonate
PEG: polyethylene glycol
LDAO: Lauryldimethylamine oxide
TFA: Trifluoroacetic acid
References:
A: Kallweit, U. et al. Rapid Comm. Mass Spec. 10, 845-849, 1996.
B: Yao, J. et al. J. Am. Soc. Mass Spectrom 9, 805-813, 1998
C: Coligan, J. E. et al. In Current Protocols in Protein Science 2, unit 16.2
D: Gevaert, K. et al. ABRF web publication, 1998.
E: Ogorzalek, et al., Protein Science 3, 1975-1983.
F. Kay, I. and Mallet, A.I. Rapid Comm. Mass Spec. 7, 744-746, 1993.
G: Funk et al. Rapid Commun. Mass Spectrom. 2005; 19: 2986–2988